Supplemental Digital Content

Table S2. Performance characteristics of the LC-MS/MS method for the determination of dialkyl phosphate (DAP) metabolites in urine

| Analyte | LOD at 3.3 σa/Sb (g L-1) | LOQ at 10 σ/S (g L-1) | Linearity (g L-1) | Correlation coefficient (r) | Precision at different concentration levels (RSDc) (g L-1) |  |  |  |  |  | % Recovery (g L-1) ± SDf (n=6) |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  |  |  |  | Intra-day (Rpd) |  |  | Inter-day (Rce) |  |  |  |  |
|  |  |  |  |  | 125 | 500 | 1000 | 125 | 500 | 1000 | 500 | 1000 |
| Dimethyl phosphate (DMP) | 0.17 | 0.55 | 0.552 - 1000 | 0.999 | 4 | 6 | 7 | 10 | 11 | 15 | 77 ± 2 | 78 ± 2 |
| Diethyl phosphate (DEP) | 0.02 | 0.07 | 0.071 - 1000 | 0.998 | 5 | 4 | 3 | 12 | 11 | 13 | 101 ± 2 | 113 ± 2 |
| Diethylthio phosphate (DETP) | 0.02 | 0.06 | 0.061 - 1000 | 0.999 | 7 | 5 | 8 | 13 | 10 | 8 | 85 ± 7 | 90 ± 7 |
| Diethyl-dithio phosphate (DEDTP) | 0.12 | 0.39 | 0.386 - 1000 | 0.999 | 6 | 8 | 5 | 10 | 12 | 15 | 79 ± 3 | 100 ± 5 |

astandard deviation of the lowest concentration response

bslope of the calibration curve

crelative standard deviation

drepeatability (n=5)

ereproducibility (n=5)

fstandard deviation