Supplementary Data

Radio-HPLC Method Validation

Identification of [F-18]PSMA-1007, PSMA-1007 and OH-PSMA-1007

For this purpose, [F-18]PSMA-1007 and 0.01mg/mL PSMA-1007 sample were examined 3 times. Different between retention time of labeled [F-18]PSMA-1007 and retention time of PSMA-1007 expressed in minute and percentage of deviation were calculated.

Table 1: Differential retention times between the chemical and radiochemical.

| tR [F-18]PSMA-1007, min | tR PSMA-1007, min | tR deviation, min | % deviation |
| --- | --- | --- | --- |
| 6.33 | 6.27 | 0.06 | 0.95 |
| 6.33 | 6.25 | 0.08 | 1.26 |
| 6.33 | 6.27 | 0.06 | 0.95 |
| Avg: 6.33 | Avg: 6.26 | Avg: 0.07 | Avg: 1.05 |

Linearity

The linearity determination was conducted in triplicate using the following concentrations: 0.03 mg/mL, 0.02 mg/mL, 0.01 mg/mL, and 0.005 mg/mL for the PSMA-1007 standard reference; and 0.15 mg/mL, 0.1 mg/mL, 0.05 mg/mL, and 0.025 mg/mL for the impurity standard reference OHPSMA-1007. For other impurities, linearity testing of the trifluoroacetate salt was performed with concentrations of 0.025 mg/mL, 0.05 mg/mL, 0.1 mg/mL, and 0.15 mg/mL.

| PSMA-1007 Concentration | 0.03 mg/ml | 0.02 mg/ml | 0.01 mg/ml | 0.005 mg/ml | Y-intercept |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Area Under curve |  |  |  |  |  |  |  |
| Run #1 | 2364 | 1561 | 782 | 384 | -11.9 |  |  |
| Run #2 | 2320 | 1587 | 776 | 398 | 13.4 |  |  |
| Run #3 | 2397 | 1596 | 783 | 390 | -15.1 |  |  |
| Average | 2360.3 | 1581.4 | 780.5 | 390.8 |  |  |  |
| st. deviation | 31.295 | 15.003 | 3.071 | 5.837 |  |  |  |
| RSD | 1.33 | 0.95 | 0.39 | 1.49 | Should be less than 2% |  |
| Repeatability | 1.04 | Should be less than 2% |  |  |  |  |  |
| Linearity (R2) | 1.000 | Should be larger than 0.995 |  |  |  |  |  |

Figure 1: Calibration curve of PSMA-1007 reference standard using the rHPLC method

Precision

The precision of the chromatographic method was evaluated by injecting the PSMA-1007 reference standard at a concentration of 0.01 mg/mL in triplicate. The peak area and retention time were measured, and the relative standard deviation (RSD) was calculated. An RSD of less than 2% was required to meet the acceptance criteria. The results are summarized in Table 2 and Table 3.

Table 2: Area under curve (AUC) of 0.01mg/mL PSMA-1007 for precision analysis

| Area Under curve | 0.01 (mg/mL) |  |  |  |
| --- | --- | --- | --- | --- |
| Run #1 | 782 |  |  |  |
| Run #2 | 776 |  |  |  |
| Run #3 | 783 |  |  |  |
| Average | 780.5 |  |  |  |
| standard deviation | 3.071 |  |  |  |
| RSD | 0.39% | Should be less than 2% |  |  |

Table 3: Retention time (tR) of 0.01mg/mL PSMA-1007 for precision analysis

| Retention Time | 0.01 (mg/mL) |  |  |  |
| --- | --- | --- | --- | --- |
| Run #1 | 6.27 |  |  |  |
| Run #2 | 6.250 |  |  |  |
| Run #3 | 6.270 |  |  |  |
| Average | 6.263 |  |  |  |
| Standard deviation | 0.009 |  |  |  |
| RSD | 0.15% | Should be less than 2% |  |  |

Accuracy

Four different concentration of the standard reference PSMA-1007 were prepared; 0.03mg/ml, 0.02 mg/ml, 0.01 mg/ml and 0.005 mg/ml. Each concentration of the reference then was injected in triplicate. The percentage recovery was calculated, with the acceptable recovery range set at within 90-110%.

Table 4: Percentage (%) of recovery for accuracy analysis.

| No of run | PSMA-1007 Concentration (mg/ml) | Area Under curve | Average | % of recovery | Average % of recovery |
| --- | --- | --- | --- | --- | --- |
| Run #1 | 0.03 | 2364 | 2360 | 100.14 | 100.0 |
| Run #2 |  | 2320 |  | 98.31 |  |
| Run #3 |  | 2397 |  | 101.55 |  |
| Run #1 | 0.02 | 1561 | 1581 | 98.70 | 100.0 |
| Run #2 |  | 1587 |  | 100.37 |  |
| Run #3 |  | 1596 |  | 100.93 |  |
| Run #1 | 0.01 | 782 | 780 | 100.17 | 100.0 |
| Run #2 |  | 776 |  | 99.46 |  |
| Run #3 |  | 783 |  | 100.37 |  |
| Run #1 | 0.005 | 384 | 391 | 98.23 | 100.0 |
| Run #2 |  | 398 |  | 101.88 |  |
| Run #3 |  | 390 |  | 99.89 |  |

Repeatability

Repeatability analysis was conducted by injecting the PSMA-1007 reference standard at four different concentrations in triplicate. The repeatability of the test was assessed by calculating the RSD for both the area under the curve and retention time.

Table 5: Area under curve (AUC) of PSMA-1007 for repeatability analysis.

| PSMA-1007 Concentration | 0.03mg/ml | 0.02mg/ml | 0.01mg/ml | 0.005mg/ml |  |
| --- | --- | --- | --- | --- | --- |
| Run #1 | 2364 | 1561 | 782 | 384 |  |
| Run #2 | 2320 | 1587 | 776 | 398 |  |
| Run #3 | 2397 | 1596 | 783 | 390 |  |
| Average | 2360.3 | 1581.4 | 780.5 | 390.8 |  |
| Standard deviation | 31.295 | 15.003 | 3.071 | 5.837 |  |
| RSD | 1.33 | 0.95 | 0.39 | 1.49 |  |
| Repeatability | 1.04 (Should be less than 2%) |  |  |  |

Table 6: Retention time (tR) of PSMA-1007 for repeatability analysis.

| PSMA-1007 Concentration | 0.03mg/ml | 0.02mg/ml | 0.01mg/ml | 0.005mg/ml |  |
| --- | --- | --- | --- | --- | --- |
| Run #1 | 6.250 | 6.25 | 6.27 | 6.27 |  |
| Run #2 | 6.270 | 6.250 | 6.250 | 6.250 |  |
| Run #3 | 6.250 | 6.250 | 6.270 | 6.250 |  |
| Average | 6.257 | 6.250 | 6.263 | 6.257 |  |
| Standard deviation | 0.009 | 0.000 | 0.009 | 0.009 |  |
| RSD | 0.15 | 0.00 | 0.15 | 0.15 |  |
| Repeatability | 0.11 (Should be less than 2%) |  |  |  |

Limit of Detection (LOD) and Limit of Quantification (LOQ)

Table 7: Limit of Detection (LOD) and Limit of Quantification (LOQ) for PSMA-1007, OHPSMA-1007 trifluoroacetate salt

|  | LOD (mg/ml) | LOQ (mg/ml) |
| --- | --- | --- |
| PSMA-1007 | 0.00065 | 0.00198 |
| OHPSMA-1007 | 0.00377 | 0.01143 |
| Trifluoroacetate salt | 0.00331 | 0.01004 |

Specificity test

For the specificity of rHPLC, 20 µL of each reference standard; 0.01 mg/mL PSMA-1007,  0.05 mg/mL OHPSMA-1007, 0.05 mg/mL Trifluoroacetate salt,  and 20 µL of the final product [18F]PSMA-1007 were injected in triplicate. The retention times of all solutions were then calculated and compared.

Table 8: Specificity test of the reference standard PSMA-1007, OHPSMA-1007, Trifluoroacetate salt and final product [F-18]PSMA-1007.

| Retention time (tR) | PSMA-1007 | OHPSMA-1007 | Trifluoroacetate salt | [F-18]PSMA-1007 |
| --- | --- | --- | --- | --- |
| Run #1 | 6.27 | 4.8 | 4.40 | 6.33 |
| Run #2 | 6.25 | 4.830 | 4.40 | 6.33 |
| Run #3 | 6.27 | 4.800 | 4.40 | 6.33 |
| Average | 6.26 | 4.810 | 4.400 | 6.33 |
| Standard Deviation | 0.01 | 0.014 | 0.000 | 0.00 |
| RSD | 0.15 | 0.29 | 0.00 | 0.00 |

System Suitability Testing (SST) of HPLC

The system suitability test (SST) is widely used in laboratories to verify that the complete analytical system is appropriate for its intended application. For rHPLC, the SST ensures the system’s suitability and effectiveness before use. For [F-18]PSMA-1007, as per the European Pharmacopoeia, the SST was performed by measuring the resolution between two peaks corresponding to the impurities defluorotrimethylaminium-PSMA-1007 trifluoroacetate (trifluoroacetate salt) and defluorohydroxy-PSMA-1007 (OHPSMA-1007). The required minimum resolution between these peaks is 2. Based on three injections of the mixture, the resolutions obtained were 3.05, 4.01, and 3.80, respectively.

Results summary

The chemical purity test of PSMA-1007 was validated following ICH Q2 (R2) guidelines. The validation parameters, specifications, and results are summarized and presented in Table 1. All obtained results met the required criteria.

| Inspected parameter | Limit | Result | Pass/Fail |
| --- | --- | --- | --- |
| Linearity | R2 >0.990 | 1.0 | Pass |
| Precision (RSD) |  |  |  |
| RSD (Rt) | <2.0 % | 0.39% | Pass |
| RSD (AUC) | <2.0 % | 0.15% | Pass |
| Accuracy | 90%-110% | 100% | Pass |
| Repeatability |  |  |  |
| RSD (Rt) | <2.0 % | 1.04% | Pass |
| RSD (AUC) | <2.0 % | 0.11% | Pass |
| Specificity | <2.0 % | 0.01% | Pass |
| LOD | (σ x 3.3)/Slope | 0.00065 | Pass |
| LOQ | (σ x 10)/Slope | 0.00198 | Pass |

Table 1: Parameters, specifications and results of chemical purity validation using rHPLC