Supplemental Information 1.

Bioanalysis of Solithromycin

Solithromycin (CEM-101) and its two circulating metabolites (i.e., N-acetyl solithromycin and CEM-214) were measured in liquid plasma samples (LPS), whereas only the parent compound was measured in dried blood spot (DBS) samples. LPS and DBS samples were analyzed in a central laboratory (MicroConstants, Inc., San Diego, CA, USA; human plasma) using validated liquid chromatography-tandem mass spectrometry methods for both matrices. Accuracy and precision observations were within the U.S. Food and Drug Administration bioanalytical assay validation criteria for both methods.

LPS samples of solithromycin were quantified using an internal standard of N-acetyl CEM-101-d6 and the anticoagulant potassium ethylenediaminetetraacetic acid (K2EDTA). Samples were diluted with water and extracted using solid-phase extraction (SPE) well plates. Analysis of the samples was performed using a reversed-phase high-performance liquid chromatography (HPLC) using a CN 100Å column (150 × 2 mm, 5 µm; Phenomenex Luna, Torrance, CA, USA) maintained at 25 °C. The SPE activation and elution solutions were composed of a 50:50 (v:v) mixture of 0.1% formic acid in methanol:acetonitrile, and the SPE wash solution was composed of a 50:50 (v:v) mixture of water:methanol. The mobile phase was nebulized using heated nitrogen in a Z-spray source/interface set to electrospray positive ionization mode. The mobile phase was composed of two solvents: solvent A (35%) comprising 20 mM ammonium formate, 0.2% formic acid, and 0.0002% citric acid in water; and solvent B (65%) comprising 0.1% formic acid in methanol:acetonitrile (50:50, v/v). The HPLC conditions were as follows: an injection volume of 20 μL, a sample temperature of 5 °C, a flow rate of 0.3 mL/min, and an analysis time of approximately 4 min. A Waters ACQUITY HPLC system (Milford, MA, USA) was used for method validation. The mass spectrometer conditions were as follows: a mass transition of 845.7 > 158.3 (for solithromycin), a cone voltage of 10 V, collision energy of 38 eV, and a dwell time of 0.2 s. A Waters Micromass Quattro Premier tandem quadrupole mass spectrometer (Milford, MA, USA) was used for method validation. The lower limit of quantitation was 0.01 μg/mL for solithromycin. The calibration range was 0.01 to 20 μg/mL for solithromycin.

DBS samples of solithromycin were quantified using an internal standard of CEM-101-d3 and the anticoagulant, K2EDTA. DBS samples were analyzed using reversed-phase HPLC with a Waters XSELECT CSH C18 column (Milford, MA, USA) maintained at 40 °C. The desorption solution was composed of 8.5:1.5:0.1 v:v:v mixture of dimethylformamide:water:acetic acid, and the injection solution was composed of 65:35 v:v mixture of 20 mM ammonium formate, 0.2% formic acid, 0.0002% citric acid in water: 0.1% formic acid in acetonitrile. Heated nitrogen in a turbo spray (AB Sciex, Framingham, MA, USA) source/interface was used to nebulize the mobile phase, and the electrospray positive mode was used to detect ionized compounds. The mobile phase was composed of two solvents: solvent A (64%) comprising 20 mM ammonium formate, 0.2% formic acid, and 0.0002% citric acid in water; and solvent B (36%) comprising 0.1% formic acid in acetonitrile. The HPLC conditions were as follows: an injection volume of 20 μL, a sample temperature of 5 °C, a flow rate of 450 μL/min, and a run time of approximately 3.5 min. An Agilent 1100 series binary gradient HPLC system (Santa Clara, CA, USA) was used for method validation. The mass spectrometer conditions were as follows: a mass transition of 423.3 >158.1, a declustering potential of 61 V, an entrance potential of 10 V, collision energy of 23 eV, a collision cell exit potential of 12 V, and a dwell time of 300 ms. An AB Sciex tandem quadrupole mass spectrometer API 4000 was used for method validation. The calibration range for the method was 0.01 to 20 μg/mL. Three-millimeter punches of human DBS were used for extraction.